시 방 서

| 제품명 | UTR-100 RTP SYSTEM |
| --- | --- |
| 용도설명 | - Rapid annealing (급속 열처리) - Diffusion/Implant anneal process (SD, well anneal) - 나노 디바이스용 thin gate oxidation - Dielectric film densification - Silicidation (TiSi, CoSi 등) - 대기압 공정 및 저압 공정 가능 - 넓은 범위의 온도 대역 운용 가능 |
| 상세사양 | UTR-100 RTP SYSTEM <개요> ◎ Heating Performance - Temperature range : RT~1000℃ - Ramp up rate : < Max. 20℃/sec - Ramp down rate : < 5.0℃/sec - Temp. uniformity : ± 1% (@1,000℃) - Temp. overshoot : < 5.0℃ ◎ Substrate size & load capacity : 4inch wafer - 1 wafer/run ◎ Halogen Lamp ◎ Quartz substrate holder ◎ Manually top cover open/close (gas spring) ◎ Rotary pump ◎ MFC: Ar, O2, N2 ◎ Manual control (switch panel) <기본사양> 1. Process Chamber Module 1) Process chamber - Chamber : Au coated on SUS chamber inside surface (Process chamber is one body SUS block) - Reflector : Au coated brass body - Quartz support pin for support graphite disc(for improve uniformity) - Fused quartz window - Chamber purge & vent - Gas injection port and pumping port is symmetry structure - T/C Feed-through with ceramic isolation - Round shape T/C quartz - Bowl type process chamber - Water flow sensor for heater safety - 2- stage chamber support plate - Movable align support block - Chamber up/down operation by support block - Unique designed flow path for lamp - Flange for pyrometer at chamber center - SiC ring for improved uniformity 2) Substrate holder unit - Quartz substrate holder : 4inch wafer compatible - Wafer loading by open the top plate - Ceramic electrode for halogen lamp - Unique designed flow path for lamp - Absorber for process chamber 3) Heating source unit - Tungsten halogen lamp : 1250W/ea - Lamp holder made by water cooled Au coated on brass - Hemisphere type each lamp reflector - Light guide made by Au coated on Al body - Water cooling - Water flow sensor for heater safety - Temperature controller editable 4 profiles and can save 40 steps 2. Vacuum Module 1) Vacuum pump - Rotary pump : 180 L/min - Ultimate pressure : 5×10-3 Torr 2) Pressure gauge - 1/4inch VCR connection type convectron gauge - Rugged all-metal package is highly immune to RF interference. - Digital type pressure readout & cable kit 3) Vacuum valves & lines - Pneumatic type angle valve - Pneumatic Soft pumping valve & pumping line - Auto vent line - Stainless steel hard line and flexible bellows line 3. Gas Delivery Module 1) Used gases & flow control - Process : O2, N2, Ar (MFC + 3채널 MFC controller) - Purge & vent : N2 (Metering valve) 2) Gas valves, gas line & water lines - Pneumatically operated diaphragm valve - All gas lines & valves using VCR fitting - Metering valve for N2 purge & vent - The gas line is helium-leak tested to 10-8 Torr·L/sec - Stainless steel hard line and flexible bellows line - All water line using Teflon tube - All gas line & water line fitting using SWAGELOK 4. Control Module 1) System control - System is controlled manually 2) Main Frame with electrical contact mechanism - Electrical power drive panel (ON/OFF/Emergency switch) - Keypad control panel - MFC controller panel - Temperature controller panel - Vacuum gauge controller panel - Used circular type connector 5. Frame Module - System frame is mild steel - Black & White colored panels & frame covers - Movable Keypad panel - Sliding type MFC module - “L” shape one body frame - Used circuit breaker handle - Curved shape one body upper & front & side frame [기타] ◎ 납기 : 계약후 6주 ◎ 납품조건 : 현장 설치 및 시험운전 ◎ 하자보증 : 검수 후 1년 [일반사양] 전원 220VAC 3상 75A 외관 950(W)×600(D)×550(H)mm |
| 표준액세서리 | Power cable, User manual, Rubber feet |